## Slide 1
Project Flexible, Biocompatible, and Wirelessly Powered Implantable Optoelectronics | What I did last four weeks? Design final version of the circuit. Design antenna Write abstract for conference paper Developed conceptual model of the device.
Problem being faced? ANTENNA | What to do in next two weeks? Find a solution for the antenna. Final etching of the device.
Gabriel Galeote Checa

## Slide 2: Bioimplantable Device
Antenna Designed
1. Current work
Parameters:
Number of turns = 5
wire width = 0.5 mm
Spacing = 0.5 mm
Outer Diameter = 22.2 mm
Inner Diameter = 12.3 mm

## Slide 3: Bioimplantable Device
Planify tests for the final circuit:
  Which tests to run?
  What to measure?
  Which environment to use?
Planify testing phase
Main points of the implant:
  Modularity
  Long shank (or variable shank)
  Flexibility
  Wireless Power Transmission (WPT)
Prepare conference Paper
1. Current work

## Slide 4: Bioimplantable Device
Resonant frequency  around 8 MHz
Using 10 Vpp (5V amplitude)
Tested in diferent angles and curvatures
2. Experimental Results

## Slide 5: Bioimplantable Device
Not tunned antenna
2. Experimental Results

## Slide 6: Bioimplantable Device
The antena firstly needs to be coupled with the primary coil. Then, we can move the antena to a certain reange and rotate it in XYZ axis.
2. Experimental Results
The power on the led seems to remain stable thanks due to the LDO Voltage regulator. However, I need a way to measure the power while I am moving the antenna in different angles and distances to corroborate that the performance is correct despite the geometrical factor

## Slide 7: Bioimplantable Device
Order primary coil (13.56 MHz)
Tune antena with capacitance
  In case that in two weeks we are not able to tune the capacitance, we will use a comercial secondary coil (13.56 MHz) for the testing
Start Testing Phase
  Rotations over XYZ in the magnetic field.
  Propagation distance.
  Performance in saline solution.
  Using an material with some electromagnetic mitigation between antenna and implant.
  The effect on the emitted light using different resistances.
3. Next steps

## Slide 8: Bioimplantable Device
Order primary coil (13.56 MHz)
Tune antena with capacitance
  In case that in two weeks we are not able to tune the capacitance, we will use a comercial secondary coil (13.56 MHz) for the testing
Start Testing Phase
  Rotations over XYZ in the magnetic field.
  Propagation distance.
  Performance in saline solution.
  Using an material with some electromagnetic mitigation between antenna and implant.
  The effect on the emitted light using different resistances.
3. Next steps

## Slide 9: Bioimplantable Device
Should I change the frequency of the implant if Welektronik cannot provide us a transmitter coil for 13.56 MHz?
Should I just use the 8 MHz antena that I am already using for the tests or should I wait until I have the final design?
Any recommendations of how to test the power of the antena transmission?
Feedback

## Slide 10: Bioimplantable Device
Thank for the attention